## Ref (hidden)
FLUORO | Exam Type | Exceed Dose? | CT SCAN | Gender | AUDIT (LAMPIRAN 7) | Lampiran 8A | Image Quality Marking | Lampiran 9
| Cardiac | Yes |  | Male | Conventional Processor | FFDM | 1 | P
| Non-Cardiac | No |  | Female | CR System | CR System | 0 | G
| ERCP |  |  |  | DR System | Conventional |  | M
| GI Lower | [6 empty] | I
GI Upper
MCU
Cerebral
ESWL
PTCA
Vascular
OTHERS
Nephrostomy
PTBD
Sinogram
Anal Fistulogram
Ascending Urethrogram
Lower Limb Angio
Cystography
Renal Embolization

## Borang 2A
Seksyen Khidmat Teknikal BKRP (2025)_Pindaan QAP Radiologi 2023
BORANG 2A: REPORT PERCENTAGE OF RETAKES FOR DIGITAL RADIOGRAPHY
FACILITY :
MACHINE MODEL/YEAR OF MACHINE :
YEAR OF INSTALLATION :
REPORT YEAR :
STANDARD : |  | ≤5%
TYPE OF ERROR |  | MONTH | [11 empty] | TOTAL | PERCENTAGE OF ERROR (%)
|  | Jan. | Feb. | Mar. | Apr. | May | June | July | Aug. | Sept. | Oct. | Nov. | Dec.
HUMAN FAULTS | 1. Over Exposure/ High Index | [12 empty] | 0 | 0%
| 2. Under Exposure/ Low Index | [12 empty] | 0 | 0%
| 3. Double Exposure | [12 empty] | 0 | 0%
| 4. Wrong Technique | [12 empty] | 0 | 0%
| 5. Wrong Patient/ Exam | [12 empty] | 0 | 0%
| 6. No Primary/ Wrong Marker | [12 empty] | 0 | 0%
| 7. Collimation Error | [12 empty] | 0 | 0%
| 8. Patient Movement | [12 empty] | 0 | 0%
| 9. Patient Related Artifact | [12 empty] | 0 | 0%
EQU IPM ENT | 10. Equipment Fault (X-ray Tube/ Grid/ Bucky) | [12 empty] | 0 | 0%
| 11. Detector/Imaging Plate | [12 empty] | 0 | 0%
| 12. Image Artifact | [12 empty] | 0 | 0%
| 13. Processing Fault | [12 empty] | 0 | 0%
OTHERS | 14. Miscellaneous. Please specify:…………….. | [12 empty] | 0 | 0%
Total number of error |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Total number of retakes exposure (A) | [13 empty] | 0
Total number of images (B) | [13 empty] | 0
Percentage (%) (A÷B) x 100% | [13 empty] | 0% | bar: MONTH Jan. to MONTH Dec.
COMMENT :
CORRECTIVE ACTION :
REPORTED BY : |  | MONTH Jan.: VERIFIED BY :
…..................................... |  | MONTH Jan.: …......................................
Name : |  | MONTH Jan.: Name of Supervisor* :
Position : |  | MONTH Jan.: Position :
Date : |  | MONTH Jan.: Date :
*Supervisor: Radiologist, Ketua Juru X-ray atau Orang Yang Bertanggungjawab
Note on B12: Explaination: 1. Images that demonstrate high density (too dark). The bony trabecular pattern and structure of soft tissue in the area of interest are not well visualised. 2. Images that over penetrated – images demonstrate high density. Bony cortical outline of thin and thick structure cannot be seen in the single radiograph.
Note on B13: Explaination: 1. Images that demonstrate low density (too light). The bony trabecular pattern and structure of soft tissue in the area of interest are not well visualised. 2. Images that under penetrated – images demonstrate low density. Bony cortical outline of thin and thick structure cannot be seen in the single radiograph.
Note on B14: Explaination: 1. There is more than one image. 2. Double exposure results when the receptor/radiographic cassette is exposed more than 1 time, and the images appear superimposed onto each other.
Note on B15: Explaination: 1. Wrong positioning – patient incorrectly position/place, image not true lateral / not true AP 2. Wrong alignment of x ray tube 3. Wrong angulation of x-ray tube 4. Wrong source image distance (SID) 5. Wrong centering point 6. Wrong Region of interest (ROI)
Note on B16: Explaination: 1. Wrong patient /wrong patient identification 2. Wrong exam/ wrong request by doctor/clinician (wrong body part/wrong view) 3. Wrong order (IT Hospital) 4. Wrong part of examination done
Note on B17: Explaination: 1. No primary marker 2. Wrong marker 3. Marker outside collimation 4. Marker superimposed with Region of Interest and etc
Note on B18: Explaination: Collimation too small/too big or incorrect which leads to examination part requested to be cut off.
Note on B19: Explaination: 1. Image blurring due to patient movement/condition (Uncooperative, restless, Short of breath, rapid breathing (tachypnoea), on ventilator and etc. 2. Anatomical part requested has been cut off because patient move before or during exposure
Note on B20: Explaination: 1. Radio opaque material attached/on the patient that include in the radiograph such as underwired bra, key in the pocket, jewellery, coin, belt, hair clip and etc. 2. Artifact due patient clothing, patient hair and etc.
Note on B21: Explaination: 1. Light Beam Diaphragm fault causing collimation error/ cone cut 2. Alignment problem 3. Electronic failure 4. Collimation knob problem 5. Bucky/ Grid error (grid cut-off) 6. AEC malfunction/problem 7. Electrical fault 8. Network error and etc.
Note on B22: Explaination: 1. Mechanical damage & electronic dysfunction. 2. Dirty IP plate/detectort. 3. Ghosting image due to improper erasing grey scale of detector problem. 4. Detector faulty and etc.
Note on B23: Explaination: 1. Line cause from dirt collected in CR reader. 2. Damage laser beam head in CR reader appears as multiple linear white line and etc.
Note on B24: Explaination: 1. Network error. 2. Mechanical issues/problems from CR reader. 3. IP plate stuck inside reader. 4. Improper scanning / error during image reading. 5. Sudden shutdown (electrical failure) and etc.
Note on B25: Explaination: Rejected by doctor/clinician due to insufficient diagnostic preference such as ryles tube insertion/ implant monitoring (Orthopedic cases / pediatric cases) and etc.

## Borang 7A
Seksyen Khidmat Teknikal BKRP (2025)_Pindaan QAP Radiologi 2023
BORANG 7A: AUDIT FORM A IMAGE QUALITY ASSESSMENT : CHEST RADIOGRAPH FOR ADULT
FACILITY :
TYPE OF MACHINE/YEAR OF MACHINE :
AUDIT YEAR :
TYPE OF IMAGE PROCESSOR : (Choose one only. Pick the main processor used)
MARKING FOR EACH CATEGORY: (All criteria have to be fulfilled) a) 1 : YES b) 0 : NO
NUM. | CRITERIA | RADIOGRAPH/IMAGE
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
PART 1.0 - IMAGE ANNOTATION
1.1 | Patient identification
1.2 | Date of Examination
1.3 | Name of Clinic/Hospital
1.4 | Primary anatomical marker
PART 2.0 - PATIENT POSITIONING
2.1 | Symmetrical positioning of thorax Stemoclavicular joints should be equidistant.
2.2 | Medial border of the scapulae to be outside of the lung fields
PART 3.0 - IMAGE COLLIMATION
3.1 | Collimation of image a) The upper border of the illuminated fields should be slightly above the shoulders bilaterally to include both apices without superimposition of the chin and the lower border down to the level of T12/ L2 to include the diaphragm
b) The lateral border to include the rib cage and part of shoulder joints.
PART 4.0 - IMAGE QUALITY ASSESSMENT
4.1 | Contrast and exposure are adequate a) Able to visualize retro cardiac area.
b) Linear and recticular details seen out to the lung periphery.
C) Able to visualize vertebral body T1 to T4.
4.2 | The lungs are well inflated Able to visualize either six ribs anteriorly or ten ribs posteriorly.
4.3 | Sharp visualization of normal chest anatomy Visually sharp : a) Trachea and proximal bronchi
b) Borders of the heart and aorta
C) Diaphragm and costo-phrenic angles
TOTAL MARKS / 15 *For each radiograph, minimum 12 score to pass audit |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SCORING (1 or 0) |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
PERCENTAGE (%) | bar: RADIOGRAPH/IMAGE 1 to RADIOGRAPH/IMAGE 15
**Adapted from European Guideline on Quality Criteria for Diagnostic Radiographic Images, EUR 16260 EN, 1996
Mandatory Criteria : PART 1.0 - IMAGE ANNOTATION
If any of Part 1.0 is not fulfilled, it will result in the automatic failure of that radiograph.
Minimum numbers of radiographs to pass are 80% (8 out of 10( for GP’s and Health Clinics) or 12 out of 15 (for All Hospitals/ Medical Institutions) radiographs )
OVERALL COMMENT :
AUDITED BY : |  | RADIOGRAPH/IMAGE 1: VERIFIED BY :
…..................................... |  | RADIOGRAPH/IMAGE 1: …......................................
Name : |  | RADIOGRAPH/IMAGE 1: Name of Radiologist :
Position :
Date : |  | RADIOGRAPH/IMAGE 1: Date :
Note on C7: Image Processor: Sila pilih jenis 'image processor' dari dropdown
Note on C12: Sistem Permarkahan No. 1-15: Pilih 1(YES) untuk patuh, pilih 0 (NO) untuk tak patuh

## Borang 7B
Seksyen Khidmat Teknikal BKRP (2025)_Pindaan QAP Radiologi 2023
BORANG 7B: AUDIT FORM B IMAGE QUALITY ASSESSMENT : EXTREMITIES RADIOGRAPH FOR ADULT
FACILITY :
TYPE OF MACHINE/YEAR OF MACHINE :
AUDIT YEAR :
TYPE OF IMAGE PROCESSOR : (Choose one only. Pick the main processor used)
MARKING FOR EACH CATEGORY: (All criteria have to be fulfilled) a) 1 : YES b) 0 : NO
NUM. | CRITERIA | RADIOGRAPH/IMAGE
|  |  | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
PART 1.0 - IMAGE ANNOTATION
1.1 | Patient identification
1.2 | Date of Examination
1.3 | Name of Clinic/Hospital
1.4 | Primary anatomical marker
PART 2.0 - PATIENT POSITIONING
2.1 | Correct region of interest included
2.2 | Two appropriate projections (AP/ Lateral/ Oblique)
PART 3.0 - X-RAY BEAM COLLIMATION
3.1 | Collimation of image
PART 4.0 - IMAGE QUALITY ASSESSMENT
4.1 | Optimum exposure a) Sharp visualization of trabeculae and cortex
b) Appropriate visualization of soft tissues
4.2 | Proximal and distal joints to be included
TOTAL MARKS / 10 *For each radiograph, minimum 8 score to pass audit |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
SCORING (1 or 0) |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
PERCENTAGE (%) |  | 0%
**Adapted from European Guideline on Quality Criteria for Diagnostic Radiographic Images, EUR 16260 EN, 1996
Mandatory Criteria : PART 1.0 - IMAGE ANNOTATION
If any of Part 1.0 is not fulfilled, it will result in the automatic failure of that radiograph.
Minimum numbers of radiographs to pass are 80% (8 out of 10( for GP’s and Health Clinics) or 12 out of 15 (for All Hospitals/ Medical Institutions) radiographs )
OVERALL COMMENT :
AUDITED BY : |  | VERIFIED BY :
…..................................... |  | …......................................
Name : |  | Name of Radiologist :
Position :
Date : |  | Date :
Note on C7: Image Processor: Sila pilih jenis 'image processor' dari dropdown
Note on C12: Sistem Permarkahan 1-15: Pilih 1 (YES) untuk patuh dan 0 (NO) untuk tak patuh.